## Slide 1: One-year outcomes from the ACURATE neo2 PMCF Post-market Registry
Won-Keun Kim, MD
Kerckhoff-Klinik GmbH, Bad Nauheim, Germany
On behalf of the ACURATE neo2 Post-Market Clinical Follow-up Study Investigators
SH-1743303 AA
CAUTION: In Europe, ACURATE neo and neo2 Aortic Valve Systems are CE-marked. In the USA, ACURATE neo2 is an investigational device and restricted under federal law to investigational use only. Not available for sale.

## Slide 2: Potential conflicts of interest
Speaker's name : Won-Keun Kim
☑ I have the following potential conflicts of interest to declare:
Receipt of grants / research support: Boston Scientific
Receipt of honoraria or consultation fees: Abbott, Boston Scientific, Edwards Lifesciences, Meril, Shockwave Medical
SH-1743303 AA

## Slide 3: Why this study?
ACURATE neo2 is a 2nd-generation supra-annular valve that features improvements to further reduce PVL compared to the prior-generation ACURATE neo.
The PMCF post-market registry studied neo2 in routine clinical practice.
Outcomes at 30 days have been previously presented. Here, we report 1-year outcomes.
CE marked 2020
Self-expanding nitinol frame
Top-down deployment
Porcine pericardial leaflets in a supra-annular position
Enhanced sealing skirt
SH-1743303 AA

## Slide 4: What did we study?
Prospective multicentre single-arm post-market surveillance study in a routine clinical practice setting
Independent CEC adjudication of safety events
  Primary safety endpoint: 30-day all-cause mortality
  Additional endpoints evaluated through 1 year
    Death, stroke, bleeding, major vascular complications, hospitalization for valve-related symptoms
Independent core laboratories
  Echocardiography: MedStar Health Research Institute
    Hemodynamics and PVL at discharge, 30 days, and 1 year
  4D-CT: The University of British Columbia, Department of Radiology
    Hypo-attenuated leaflet thickening (HALT) at 30 days [primary imaging endpoint] and 1 year
SH-1743303 AA

## Slide 5: How was the study executed?
*In 4 patients, ACURATE neo2 embolized and patients were implanted with a non-study valve; these patients were followed for safety only through 30 days
†Patient survival confirmed, but no additional follow-up data collected.
30-day Clinical Follow-up Visit Completed or Death 94% (236/250)
Premature discontinuation
Withdrawal from study, n=20
Missed 1y follow-up, n=7 †
Successfully Implanted
N=246
1-year Clinical Follow-up Visit Completed or Death 89% (223/250)
TTE: 91% (224/246)
30-day Imaging Data Analyzed
CT: 83% (204/246)
TTE: 76% (186/246)
1-year Imaging Data Analyzed
CT: 62% (153/246)
Intention-To-Treat
N=250
Mean Age
80.8 ± 6.2 yrs
36.4% Male
63.6% Female
Mean STS Score
2.9% ± 2.0%
Operative Risk
(per Heart Team)
Chart: 
| High | Intermediate | Low
STS Score | 31.2 | 31.6 | 37.2
High
Low
Intermediate
Device not in correct position n=4*
SH-1743303 AA

## Slide 6: What are the essential results?
Key Safety Events | Discharge | 30 Days | 1 Year
All-cause mortality | 0.4% (1) | 0.8% (2)* | 5.1% (12)
Cardiovascular mortality | 0.4% (1) | 0.8% (2) | 3.4% (8)
All stroke | 0.4% (1) | 0.8% (2) | 3.0% (7)
Major bleeding | 2.4% (6) | 2.4% (6) | 2.8% (7)
Hospitalization for valve-related symptoms or worsening CHF | 0.0% (0) | 0.0% (0) | 1.7% (4)
Prosthetic aortic valve thrombosis | 0.4% (1)^ | 0.4% (1) | 0.9% (2)‡
Prosthetic aortic valve endocarditis | 0.0% (0) | 0.4% (1) | 1.7% (4)
Valve-related repeat procedure (BAV) | 0.4% (1)* | 0.4% (1) | 0.8% (2)†
Newly implanted permanent pacemaker | 5.2% (13) | 6.1% (15) | 7.8% (19)
*Deaths: 1 patient experienced life-threatening bleeding (admitted to secondary hospital, source of bleeding not recorded) and hemodynamic instability leading to circulatory arrest and death; 1 patient died post index procedure following valve embolization and unsuccessful TAV-in-TAV with non-study valve leading to coronary obstruction and procedural myocardial infarction.
*Surgical removal of embolized ACURATE neo2 valve ^Thrombus noted pre-discharge in non-study valve implanted subsequent to ACURATE neo2 embolization ‡Diagnosis of HALT at 30 days and 1 year, with oral anticoagulant therapy initiated shortly before 1-year visit
1° Safety Endpt
SH-1743303 AA

## Slide 7: What are the essential results?
Chart: 
| BL | Discharge | 30 Days | 1 Year
Grad | 47.6 | 9.7 | 8.6 | 7.6
EOA | 0.72 | 1.57 | 1.58 | 1.67
Mean AV Gradient (mmHg)
.
Mean AV Area (cm2)
Baseline
Discharge
8.6 ± 3.9 (n=180)
1.6 ± 0.4
(n=127)
30 Days
9.7 ± 5.4 (n=193)
1.6 ± 0.4
(n=139)
47.6 ± 14.5 (n=240)
0.7 ± 0.2
(n=237)
1 Year
7.6 ± 3.2 (n=169)
1.7 ± 0.4
(n=134)
Aortic Valve Haemodynamics, ITT Population
Paravalvular Leak, Paired Analysis (n=116)
Echocardiography Core Laboratory
Chart: 
| Discharge/7d | 30 Days | 1 Year
none/trace | 77.3585 | 81.9048 | 83.7838
mild | 22.6415 | 16.1905 | 15.3153
moderate | 0 | 1.9048 | 0.9009
severe | 0 | 0 | 0
None/Trace
Mild
Moderate
Severe
SH-1743303 AA

## Slide 8: What are the essential results?
CT Core Laboratory
1-Year Event Rate | HALT @ 30d or 1y N=72 | No HALT @ 30d and 1y N=93 | P-value
All-cause mortality | 0.0% (0) | 0.0% (0) | 0.25
CV mortality | 0.0% (0) | 0.0% (0) | 0.25
All stroke | 2.2% (2) | 2.2% (2) | 0.74
Disabling stroke | 0.0% (0) | 0.0% (0) | --
Valve thrombosis | 0.0% (0) | 0.0% (0) | 0.25
Echocardiography | n=27* | n=91 | 
AV gradient (mm Hg) | 7.1±4.3 | 7.8±2.7 | 0.36
AV area (cm2) | 1.6±0.4 | 1.7±0.4 | 0.53
*Only subjects with HALT recorded at 30 days and 1 year are included in the analysis.
HALT, Paired Analysis (n=150)
Chart: 
| 30 Days | 1 Year
None | 0.76 | 0.68
<50% severity | 0.146 | 0.206
>50% severity | 0.094 | 0.114
No HALT
HALT severity ≤50%
HALT severity >50%
1° Imaging Endpt
Change in HALT severity from 30 days  1 year
(n=135, excludes pts on OAC)
Chart: 
| Worsened | No change | Improved
change | 21 | 67 | 12
No change
Improved
Worsened
SH-1743303 AA

## Slide 9: The essentials to remember
Patients treated with ACURATE neo2 exhibited early improvement in valve haemodynamics and low rates of PVL through 1 year
  At 1 year, 99% of patients had mild or no/trace PVL (<1% had moderate PVL)
  No patients exhibited >moderate PVL at any time
1-year rates of all-cause mortality (5.1%) and stroke (3.0%) are comparable to other trials of TAVR patients at intermediate or high surgical risk
The presence of HALT:
  Did not impact patient safety at 1 year
  Did not significantly affect mean AV gradient or mean AV area
This report confirms favorable performance and safety through 1 year in patients with severe aortic stenosis treated with ACURATE neo2
SH-1743303 AA

## Slide 10
CAUTION: In Europe, ACURATE neo and neo2 Aortic Valve Systems are CE-marked. In the USA, ACURATE neo2 is an investigational device and restricted under federal law to investigational use only. Not available for sale.